SPS30 sensor

# 1. Background

This project is part of a number of projects to measure the air quality I bought a new sensor : SPS30 from Sensirion. Having worked on another sensor from them, SCD30, before the quality is normally good. The aim is to compare the output with other sensors and provide the learning

The board is relative new at the time of writing this documentation and as such a lot of preliminary documents from supplier. Actually the documentation looks to be very incorrect on a number of places that I will point to in this document.

While there is a driver available from Sensirion it is very modular, to a point where it feel to be over engineered and complex. Also I looked to be half-library as I did not find good user examples for Arduino / ESP32. Hence I wrote the driver from scratch and it will support UART communication (different serial ports) as well as I2C.

This is the first step. The next step is more examples, WIFI connection with ESP32 and compare with other sensors.

Paul van Haastrecht
paulvha@hotmail.com
# 2. Hardware connection

Depending on UART or I2C the connections are different.

## 2.1 UART communication
### 2.1.1 ESP32
Using serial port1, setting the RX-pin(25) and TX-pin(26)
Different setup can be configured in the sketch.

SPS30 pin ESP32
1 VCC -------- VUSB
2 RX --------- TX pin 26
3 TX --------- RX pin 25
4 Select (NOT CONNECTED)
5 GND -------- GND

Also successfully tested on Serial2 (default pins TX:17, RX: 16)
NO level shifter is needed as the SPS30 is TTL 5V and LVTTL 3.3V compatible

### 2.1.2 ATMEGA2560
Used SerialPort2. No need to set/change RX or TX pin
SPS30 pin ATMEGA
1 VCC -------- 5V
2 RX --------- TX2 pin 16
3 TX --------- RX2 pin 17
4 Select (NOT CONNECTED)
5 GND -------- GND

Also successfully tested on SerialPort1 and Serialport3

### 2.1.3 Failed UNO R3
Failed as I had to use softserial as there is not a separate serial port. But as the SPS30
is only working on 115K the connection failed all the time with CRC errors. (see 2.2.3)

### 2.1.4 Failed ESP8266
As the power is only 3V3 (the SPS30 needs 5V) and one has to use softserial which failed with 115K speed

## 2.2 I2C communication

As documented in the datasheet, make sure to use external 10K pull-up resistor on both the SDA and SCL lines. Otherwise the communication with the sensor will fail random. Also read chapter 4 about the expected valid values in I2C

### 2.2.1 ESP32

SPS30 pin ESP32
1 VCC -------- VUSB
2 SDA -------- SDA (pin 21)
3 SCL -------- SCL (pin 22)
4 Select ------ GND (select I2c)
5 GND ------- GND

The pull-up resistors should be to 3V3.

### 2.2.2 ATMEGA2560

SPS30 pin ATMEGA
1 VCC -------- 5V
2 SDA -------- SDA
3 SCL --------- SCL
4 Select ------- GND (select I2c)
5 GND -------- GND

### 2.2.3 UNO R3

SPS30 pin UNO
1 VCC ------- 5V
2 SDA ------- A4
3 SCL -------- A5
4 Select ------ GND (select I2c)
5 GND ------- GND

When UNO-board is detected the UART code is excluded as that does not work on UNO and will save memory. Also some buffers reduced and the call to GetErrDescription() is removed to allow enough memory. Apart from that ALL examples work with UNO as well :-)

### 2.2.4 ESP8266

SPS30 pin External ESP8266
1 VCC -------- 5V
2 SDA ----------------------- SDA
3 SCL ------------------------ SCL
4 Select ---------------------- GND (select I2c)
5 GND ------- GND -------- GND

The pull-up resistors should be to 3V3 from the ESP8266.
# 3. Software
This software has been written from scratch. The driver has no dependencies on other libraries, some examples have. It has a number of user level calls to set and get data from the SPS30 and can be used independent of the communication channel setting. The I2C communication however could have some restrictions. See chapter 4.
## 3.1 Examples overview
At the top of each sketch there is detailed information about the hardware connections and dependencies on external libraries.
### 3.1.1 Basic reading
This sketch will detect an SPS30 and read and display all the data every 3 seconds
It will allow to configure for communication channel, serial-RX and TX pins as well as debug level.

### 3.1.2 Basic reading with fan-clean
This sketch will detect an SPS30, read the current auto clean interval, optionally set new auto clean interval or force a manual fan-clean. After that it wile read and display all the data every 3 seconds.
It will allow to configure for communication channel, serial-RX and TX pins, debug level and options to set the auto clean interval or force a manual fan-clean.

### 3.1.3 Basic reading specific data & order
This sketch will detect an SPS30 and allows to set and display specific data in a specific order every 3 seconds. It will allow to configure for communication channel, serial-RX and TX pins, debug level and define the wanted data in a wanted order.

### 3.1.4 Basic reading with DS18X20
This sketch will detect an SPS30 and DS18X20 temperature sensor and read and display all the data every 3 seconds. It will allow to configure for the SPS30 communication channel, serial-RX and TX pins as well as debug level. For the DS18x20 it will allow to configure the pin and whether to display in Celsius or Fahrenheit. This sketch the download library: https://github.com/PaulStoffregen/OneWire. Code for the DS18x20 is based on the DS18x20 example in the onewire library
### 3.1.5 Basic reading with BME280
This sketch will detect an SPS30 and DS18X20 temperature sensor and read and display all the data every 3 seconds. It will allow to configure for the SPS30 communication channel, serial-RX and TX pins as well as debug level. For the BME280 it will allow to configure the I2C address, whether to display in Celsius or Fahrenheit and whether to display altitude in meter or foot. This sketch the download library: https://github.com/sparkfun/SparkFun_BME280_Arduino_Library

### 3.1.6 SPS30 plotter
This sketch will detect an SPS30, reads and displays the data every seconds The output format is aimed for the plotter (cntrl+shift+L). It will allow to configure for the SPS30 communication channel, serial-RX and TX pins as well as debug level. You can configure the data to plot (MassPMx, NumPMx or average size) as well as how of the first readings to skip to remove invalid start-up info.
### 3.1.7 SPS30 and SCD30
This sketch will detect an SPS30 and SCD30, reads and display data every 3 seconds. It will allow to configure for the SPS30 communication channel, serial-RX and TX pins as well as debug level. For the SCD30 it will allow to configure to display in Celsius or Fahrenheit. This sketch the depending on the SCD30 library: https://github.com/paulvha/scd30
NOTE: on an ESP32 together with and SCD30 the standard I2C/Wire library can not be used as that does not support clock stretching. The SoftWire library, part of SCD30 has been included to support that. See 3.3.5. Apart from that the I2C on an ESP32 initializes in a particular way and as part of that it will fail often if a device is constant sending data. Read the top of the sketch for more details

### 3.1.8 SPS30 and SCD30 and BME280
This sketch is a combination of example 5 and 7 in one sketch.
## 3.2 Overview of user level calls.

### 3.2.1 EnableDebugging(act)
Will enable or disabled driver debug messages. Valid options for level of debug
 0 : no debug message (default)
1 : sending and receiving data
2 : 1 + protocol progress

### 3.2.2 SetSerialPin( rx, tx)
Sets the RX and TX pins to use for softserial or serialport1 on the ESP32. If these communication channels are used the pins must be set before calling begin. See NOTE 2 with chapter 3.2.3

### 3.2.3 begin(Comms_port)
This will initialize the communication channel. Unlike in other sketches, you should NOT first perform a begin ( Serial1.begin()) in the sketch. That is all handled by the driver. The following options are supported :

I2C_COMMS	I2C communication
SOFTWARE_SERIAL	Arduino variants (see NOTE)
SERIALPORT	ONLY IF there is NO monitor attached
SERIALPORT1	Arduino MEGA2560,
ESP32 Thing : MUST define pins as defaults are for flash memory
SERIALPORT2	Arduino MEGA2560 and ESP32	(default is omitted)
SERIALPORT3	Arduino MEGA2560 only for now

NOTE 1 : Softserial has been left in as an option, but as the SPS30 is only working on 115K the connection will probably NOT work on any device. See chapter 3.3.4 to excluding Software Serial code.

NOTE 2: SerialPort1 is only enabled in case of MEGA2560 or ESP32 (with defined RX/TX before)
On some boards/IDE (like Particle Photon) Serial1 is available as well. To force-enable from the user sketch : set TX and RX pins BOTH to 8, before calling begin().
l
Will return ‘true’ if successful, else ‘false’.

### 3.2.4 probe()
This will try to connect the SPS30 over the requested communication channel.
Will return ‘true’ if successful, else ‘false’.

### 3.2.5 reset()
This will sent a soft-reset to the SPS-30.
Will return ‘true’ if successful, else ‘false’.

### 3.2.6 start()
This will sent a start-measurement command to the SPS-30.
Will return ‘true’ if successful, else ‘false’.

### 3.2.7 stop()
This will sent a stop-measurement command to the SPS-30.
Will return ‘true’ if successful, else ‘false’.

### 3.2.8 clean()
This will sent a manual clean fan command to the SPS-30. The fan will spin faster for 10 seconds.
The SPS30 must be in measurement mode. Will return ‘true’ if successful, else ‘false’.

### 3.2.9 GetAutoCleanInt(uint32_t *val)
This will read the current Auto Fan clean interval from the SPS30 and return that in pointer to ‘val’.
Will return ‘ERR_OK’if successful read, else on the error messages documentation in Appendix A.

### 3.2.10 SetAutoCleanInt(uint32_t val)
This will SET the Auto Fan cleaning interval on the SPS30 to the value defined in ‘val’. Will return ‘ERR_OK’if successful read, else on the error messages documentation in Appendix A.
### 3.2.11 GetErrDescription(code, *buf, len)
This will provide a description of an error code as documented in Appendix A. ‘Code’ the error code, ‘buf’ is a buffer to store the message and ‘len’ is the length of the buffer. (not supported on Uno : see chapter 3.2)
### 3.2.12 GetSerialNumber(*buf, len)
This will try to read the serial number from the device and store that in buffer ‘buf’ with a maximum number of bytes ‘len’. The maximum can be 32 bytes. Will return ‘true’ if successful, else ‘false’.
### 3.2.13 GetArticleCode(*buf, len)
This will try to read the Articlecode from the device and store that in buffer ‘buf’ with a maximum number of bytes ‘len’. The maximum can be 32 bytes. Will return ‘true’ if successful, else ‘false’.
### 3.2.14 GetProductName(*buf, len)
This will try to read the productname from the device and store that in buffer ‘buf’ with a maximum number of bytes ‘len’. The maximum can be 32 bytes. Will return ‘true’ if successful, else ‘false’.
### 3.2.15 I2C_expect()
Will return the number of valid float values to expect from reading in I2C. See chapter 4.
### 3.2.16 GetValues(struct sps_values *v)
This will try to read all the values from the SPS30 and store those in structure pointer v. Will return ‘ERR_OK’ if successful read else on the error messages documentation in Appendix A.
The structure is defined in sps30.h, for details read the datasheet
typedef struct sps_values
{
    float MassPM1; // Mass Concentration PM1.0 [μg/m3]
    float MassPM2; // Mass Concentration PM2.5 [μg/m3]
    float MassPM4; // Mass Concentration PM4.0 [μg/m3]
    float MassPM10; // Mass Concentration PM10 [μg/m3]
    float NumPM0; // Number Concentration PM0.5 [#/cm3]
    float NumPM1; // Number Concentration PM1.0 [#/cm3]
    float NumPM2; // Number Concentration PM2.5 [#/cm3]
    float NumPM4; / / Number Concentration PM4.0 [#/cm3]
    float NumPM10; // Number Concentration PM4.0 [#/cm3]
    float PartSize; // Typical Particle Size [μm]
};

As mentioned in chapter 4, when the communication channel is I2C, depending on the receive buffer size in wire.h, only the MassPMx may be available. The others will be zero.
### 3.2.17 GetSpecific value()
This will try to return a specific value as a float. In case of error a ‘-1’ value will be returned.
The following calls are valid :
GetMassPM1()
GetMassPM2()
GetMassPM4()
GetMassPM10()
GetNumPM0()
GetNumPM1()
GetNumPM2()
GetNumPM4()
GetNumPM10()
GetPartSize()
The driver will perform a GetValues() and report the specific requested value. If another value is requested this will be provided from cache. As such the different values are a snap-shot taken at the same time. If a value is requested that has been reported already all the values will be refreshed.
As mentioned in chapter 4, when the communication channel is I2C, depending on the receive buffer size in wire.h, only the MassPMx may be available.

## 3.3 Driver configuration

In the SPS30.h configuration settings are available

### 3.3.1 Automatic configuration
In case of an UNO or other board with small memory some changes are applied automatically.
The buffers are reduced, the error message lookup is disabled and the UART communication code is excluded.
### 3.3.2 Exclude I2C communication code
In case you do not plan to use the I2C code and want to remove that in order to reduce the memory size of the driver in sps30.h comment out the line : #define INCLUDE_I2C 1.
Compile will fail if UART communication is excluded at the same time.
### 3.3.3 Exclude UART communication code
In case you do not plan to use the I2C code and want to remove that in order to reduce the memory size of the driver in sps30.h comment out the line : #define INCLUDE_UART 1
Compile will fail if I2C communication is excluded at the same time.
### 3.3.4 Exclude Software Serial communication code
In case software serial is not available for your board / IDE (causing compile errors) OR you want to save memory in sps30.h comment out the line: #define INCLUDE_SOFTWARE_SERIAL 1.
### 3.3.5 ESP32 / SCD30 / clock stretching library
In case a sketch is to interact with both SPS30 and SCD30 on a ESP32 platform running BOTH over I2C, you MUST use the SoftWire I2C that is part of SCD30. For that you need to un-comment in sps30.h the line: // #define SOFTI2C_ESP32 1
If you platform is ESP32 and you do NOT plan to run the SPS30 over I2C, you can also apply option 3.3.2.
# 4. I2C communication

In datasheet (page 11) the disclaimer is made:

Some considerations should be made about the use of the I 2 C interface. I 2 C was originally designed to connect two chips on a PCB. When the sensor is connected to the main PCB via a cable, particular attention must be paid to electromagnetic interference and crosstalk. Use as short as possible (< 10 cm) and/or well shielded connection cables.
We recommend using the UART interface instead, whenever possible: it is more robust against electromagnetic interference, especially with long connection cables

While the above is true, the I2C implementation, which works good as long as you apply the 2 x 10K pull-up resistors as indicated in the datasheet on pag 11.

## 4.1 Number of valid values
Opposite to an earlier assumption that the SPS30 is only sending 32 bytes, it turned out that the I2C / Wire receive buffer on many boards is set a maximum value of 32. After that the boards will stop with receiving.

A size needed for each float value is 6 bytes (LSB + MSB + CRC ) ++ (LSB + MSB + CRC) To read all values an I2C receive buffer of at least 6 x 10 = 60 bytes is needed. As on many boards the default buffer size is set to 32 in Wire.h, it will only be providing 5 valid float values.

To cope with this we determine the I2C buffer-size in the sps30.h and calculate in the constructor for sps30 the expected valid bytes. If the I2C buffer size is less than 64 only the MASS values are provided.

One can check the impact by calling I2C_expect(), which will return the number of valid float values to expect from reading SPS30.

Of course you can increase the buffer (if memory size allows) for your board in Wire.h. This is NOT part of this library, but part of your IDE.

ATMEGA (UNO or 2560 .. maybe others)
I had to increase BOTH otherwise the connection would fail:
Wire/src/wire.h	#define BUFFER_LENGTH 64 // was 32
Wire/src/utility/twi.h	#define TWI_BUFFER_LENGTH 64 // was 32

esp8266
Wire/src/wire.h	#define BUFFER_LENGTH 64 // was 32

esp32
The buffer is already 128
Wire/src/wire.h	#define I2C_BUFFER_LENGTH 128
# Appendix A : Error messages
Defined in sps30.h

Hex
00 ERR_OK	“All good"
01 ERR_DATALENGTH "Wrong data length for this command (too much or little data)"
02 ERR_UNKNOWNCMD "Unknown command"
03 ERR_ACCESSRIGHT "No access right for command"
04 ERR_PARAMETER "Illegal command parameter or parameter out of allowed range"
28 ERR_OUTOFRANGE "Internal function argument out of range"
43 ERR_CMDSTATE "Command not allowed in current state"
50 ERR_TIMEOUT "No response received within timeout period"
51 ERR_PROTOCOL "Protocol error"

# Appendix B : some example output

## Example 1 : basic reading

SPS30-Example1: Basic reading. press <enter> to start
Trying to connect
Detected SPS30.
Serial number : 361D4FD3AA941A76
Product name : not available
Article code : not available
Measurement started

Hit <enter> to continue reading
Hit <enter> to continue reading
-------------Mass ----------- ------------- Number -------------- -Average-
     Concentration [μg/m3] Concentration [#/cm3] [μm]
P1.0	P2.5	P4.0	P10	P0.5	P1.0	P2.5	P4.0	P10	PartSize

13.31	16.19	18.14	18.85	107.15	103.87	106.74	107.09	0.00	0.60
12.88	14.86	16.06	16.51	103.75	101.60	103.49	103.71	0.00	0.61
12.77	14.80	16.06	16.52	102.83	100.61	102.56	102.79	0.00	0.65
12.50	14.04	14.90	15.22	100.64	99.02	100.45	100.61	0.00	0.61
12.14	13.42	14.08	14.32	97.77	96.48	97.63	97.75	0.00	0.62
11.97	13.09	13.63	13.82	96.39	95.29	96.27	96.37	0.00	0.61
11.80	12.87	13.37	13.55	95.02	93.98	94.92	95.01	0.00	0.61
11.61	12.60	13.04	13.20	93.50	92.55	93.40	93.48	0.00	0.60
11.50	12.44	12.84	12.99	92.59	91.70	92.50	92.58	0.00	0.59
11.37	12.29	12.67	12.81	91.58	90.72	91.49	91.56	0.00	0.59
11.43	12.32	12.68	12.82	92.02	91.19	91.94	92.01	0.00	0.57
11.43	12.31	12.65	12.77	92.07	91.27	92.00	92.06	0.00	0.58
11.43	12.28	12.60	12.72	92.09	91.32	92.01	92.07	0.00	0.57
11.34	12.15	12.45	12.56	91.33	90.60	91.26	91.31	0.00	0.57
11.25	12.03	12.31	12.41	90.58	89.88	90.51	90.57	0.00	0.56
11.10	11.85	12.10	12.19	89.42	88.77	89.37	89.41	0.00	0.56
11.13	11.85	12.08	12.17	89.61	88.98	89.55	89.60	0.00	0.56
11.00	11.69	11.90	11.98	88.56	87.97	88.51	88.55	0.00	0.57
10.76	11.42	11.61	11.68	86.67	86.12	86.63	86.66	0.00	0.56
10.71	11.35	11.53	11.60	86.23	85.70	86.19	86.23	0.00	0.56
10.67	11.30	11.47	11.53	85.94	85.42	85.90	85.93	0.00	0.56
10.56	11.17	11.33	11.39	85.02	84.52	84.98	85.01	0.00	0.57
10.51	11.11	11.26	11.32	84.62	84.13	84.58	84.61	0.00	0.56
10.55	11.16	11.31	11.36	85.00	84.52	84.97	85.00	0.00	0.57
10.73	11.36	11.51	11.57	86.46	85.96	86.42	86.45	0.00	0.57
11.08	11.72	11.88	11.94	89.23	88.72	89.19	89.22	0.00	0.56
11.17	11.81	11.96	12.02	89.95	89.44	89.91	89.94	0.00	0.56
11.01	11.63	11.78	11.83	88.66	88.16	88.62	88.65	0.00	0.57
10.89	11.50	11.65	11.70	87.70	87.22	87.67	87.70	0.00	0.58

## Example 5 : basic reading with BME280
SPS30-Example5: Basic reading + BME280. press <enter> to start
Trying to connect
Detected SPS30
Serial number : 361D4FD3AA941A76
Product name : not available
Article code : not available
Detected BME280
Measurement started

Hit <enter> to continue reading
================================= SPS30 ================================= ========== BME280 ============
--------------------Mass ------------------- ----------------------- Number ------------------------- -Average- Pressure Humidity Altitude Temperature
        Concentration [μg/m3] Concentration [#/cm3] [μm] [hPa]	[%] Meter	[*C]
P1.0	P2.5	P4.0	P10	P0.5	P1.0	P2.5	P4.0	P10 PrtSize

15.26	15.87	15.87	15.87	122.90	122.53	122.90	122.90	0.00	0.36	967	37.7	390.6	22.98
15.69	16.33	16.33	16.33	126.39	126.01	126.38	126.39	0.00	0.42	1005	37.8	68.8	22.93
15.90	16.54	16.54	16.54	128.06	127.67	128.05	128.06	0.00	0.44	1005	37.8	68.9	22.90
16.63	17.30	17.30	17.30	133.93	133.52	133.92	133.93	0.00	0.45	1005	37.7	67.8	22.84
17.01	17.70	17.70	17.70	137.02	136.60	137.01	137.01	0.00	0.46	1005	37.6	69.1	22.80
17.20	17.89	17.89	17.89	138.51	138.08	138.50	138.50	0.00	0.46	1005	37.5	68.5	22.77
17.20	17.89	17.89	17.89	138.51	138.09	138.50	138.51	0.00	0.47	1005	37.6	69.4	22.73
17.24	17.94	17.94	17.94	138.89	138.47	138.88	138.89	0.00	0.47	1005	37.5	68.9	22.67
17.40	18.10	18.10	18.10	140.13	139.70	140.12	140.13	0.00	0.47	1005	37.5	68.4	22.63
17.42	18.12	18.12	18.12	140.30	139.88	140.30	140.30	0.00	0.48	1005	37.3	68.6	22.58

## Example 7 : basic reading with SCD30

SPS30-Example7: Basic reading + SCD30. press <enter> to start
Trying to connect
Detected SPS30
Serial number : 1FDFF1525C39D2B1
Product name : not available
Article code : not available
SCD30 found
Serial number: 390943
Measurement started

Hit <enter> to continue reading
=================================== SPS30 ================================= ========= SCD30 ==========
---------------------Mass ----===---------- ----------------------- Number -------------------------- -Average- CO2 Humidity Temperature
     Concentration [μg/m3] Concentration [#/cm3] [μm] [ppm] [%] [*C]
P1.0	P2.5	P4.0	P10	P0.5	P1.0	P2.5	P4.0	P10 PrtSize

16.29	17.21	17.21	17.21	127.91	127.44	127.90	127.91	0.00	0.55	741	35.6	22.15
17.86	21.66	23.86	24.41	140.80	136.81	140.33	140.73	0.00	0.57	741	35.4	22.25
17.08	19.54	20.73	21.02	134.41	132.05	134.16	134.37	0.00	0.56	742	35.3	22.33
15.96	16.86	16.86	16.86	125.31	124.85	125.30	125.31	0.00	0.54	742	35.0	22.44
15.65	16.53	16.53	16.53	122.85	122.40	122.84	122.85	0.00	0.55	741	35.0	22.49
15.64	17.00	17.37	17.46	122.90	121.85	122.81	122.88	0.00	0.56	739	34.8	22.62
15.61	16.49	16.49	16.49	122.56	122.11	122.55	122.56	0.00	0.56	739	34.7	22.67
15.97	16.89	16.90	16.91	125.39	124.90	125.37	125.38	0.00	0.56	744	34.5	22.77
16.12	17.05	17.08	17.08	126.54	126.03	126.52	126.53	0.00	0.56	745	34.4	22.80
16.37	17.33	17.37	17.38	128.50	127.97	128.48	128.50	0.00	0.56	743	34.3	22.90
16.47	17.42	17.44	17.44	129.29	128.78	129.28	129.29	0.00	0.56	745	34.2	22.94
16.69	17.68	17.72	17.74	131.06	130.51	131.04	131.06	0.00	0.56	745	33.9	23.01

## Example 8 : Basic reading + SCD30 + BME280
Trying to connect
Detected SPS30
    Serial number : 1FDFF1525C39D2B1
    Product name : not available
    Article code : not available
Detected SCD30
    Serial number : 390943
Detected BME280
Measurement started

Hit <enter> to continue reading

============== SPS30 ============================== ========= SCD30 ====== ======== BME280 ===============
-------------Mass ----------- ------------- Number -------------- -Average- CO2 Humidity Temperature Pressure Humidity Altitude Temperature
     Concentration [μg/m3] Concentration [#/cm3] [μm] [ppm] [%] [*C] [hPa] [%] Meter [*C]
P1.0 P2.5 P4.0 P10 P0.5 P1.0 P2.5 P4.0 P10 PrtSize

3.04 3.63 3.97 4.05 23.96 23.34 23.88 23.94 0.00 0.60 804 32.7 24.89 1022 44.7 -74.1 21.29
2.91 3.62 4.05 4.15 22.98 22.22 22.89 22.97 0.00 0.59 802 32.6 24.89 1022 44.6 -74.0 21.29
2.57 3.05 3.31 3.38 20.25 19.76 20.19 20.24 0.00 0.60 800 32.6 24.90 1022 44.8 -74.0 21.30
2.35 2.82 3.09 3.16 18.54 18.05 18.48 18.53 0.00 0.60 799 32.5 24.87 1022 45.3 -74.8 21.30
2.18 2.50 2.66 2.69 17.13 16.81 17.09 17.12 0.00 0.56 796 32.6 24.89 1022 45.0 -74.6 21.30
2.15 2.57 2.81 2.87 16.95 16.52 16.90 16.95 0.00 0.57 797 32.7 24.87 1022 44.6 -74.4 21.30
2.01 2.34 2.52 2.56 15.82 15.48 15.78 15.81 0.00 0.57 798 32.7 24.90 1022 45.4 -74.1 21.30
1.95 2.32 2.53 2.58 15.39 15.01 15.35 15.39 0.00 0.60 798 32.8 24.89 1022 45.1 -74.5 21.31
1.84 2.21 2.42 2.47 14.48 14.10 14.44 14.48 0.00 0.58 800 32.7 24.87 1022 45.0 -74.2 21.30
1.76 2.07 2.24 2.28 13.90 13.58 13.86 13.89 0.00 0.58 801 32.7 24.89 1022 45.0 -75.1 21.30
# Appendix C: inside SPS30

Removing the top cover screw

 Carefully remove from the housing. Watch the spring at the back. Removing the 4 screws allows the top panel to be removed

Air goes in on the right hand side, through a measurement chamber and is coming out through the fan. At the bottom, the golden piece, is the led.