SPS30 sensor

# 1. Background

This project is part of a number of projects to measure the air quality. I bought a new sensor : SPS30 from Sensirion. Having worked on another sensor from them (SCD30) before the quality is normally good. The aim is to compare the output with other sensors and provide the learning.

The board was relative new, but at the time of writing this version of the documentation a better datasheet is available from supplier. (July 2020)

While there is a driver available from Sensirion it is very modular, to a point where it feel to be over engineered and complex. Also it looked to be half-library as I did not find good user examples for Arduino / ESP32. Hence I wrote the driver from scratch and it supports UART communication (different serial ports) as well as I2C.

A number of examples have been added. Tailored examples for LORA/Feather and SODAQ AFF/SARA-R410M NB-IOT are also available on my github.

In March 2020 an updated datasheet was made available. It contained information about updated SPS30 function and new functions, of which some depend on firmware revisions. The library has been updated with new functions : sleep(), wakeup(), GetVersion() and GetStatusReg(). Also some existing functions have been updated.

In May 2020, based on continued new boards with different pin-outs, as well as boards with multiple I2C channels, the option has been added to select any serial or I2C channel and pass that initialized to the library. Example12 and 13 demonstrate those options.

The library has been validated on many boards now and in June 2020 I got hold of an Arduino Due. It does not seem to support SoftSerial, so for all the friends using a Due, I made changes that you can connected it to I2C, Serial1, Serial2 and Serial3, just like it is for a Mega2560.

Thank you Sensirion !! In June 2020 they have reached out as I have 2 SPS30 ‘only’ on firmware level 1.0 and offered me SPS30 with revision 2.2. I could now test the code and made a number of changes for the new functions : sleep(), wakeup() and GetStatusReg(). I have changed example11 to demonstrate reading the status register only and added example14 to demonstrate the usage of Sleep and wakeup().

August 2020. I obtained connectors for all 3 SPS30 and added Example20 which can run up to 5 SPS30 devices in parallel on a single Arduino Due, 4 devices on Mega2560 or ESP32. See chapter 9 for explanations.

October 2020 : added a Sparkfun LCD to display the results (with optional warning limit)

Paul van Haastrecht
paulvha@hotmail.com

# 2. Hardware connection

Depending on UART or I2C the connections are different. Has been tested by others on many other platforms including SAMx1D.

## 2.0 Select any communication
From the start I had decided to take an embedded communication channel setup. This made it much easier for the user to select and have the right setup for Serial or I2C. It was initially tested on an UNO, MEGA, ESP32, ESP8266 and works well.
Overtime code has been adjusted to support many more boards with different pin-outs based on user request and feedback. The Arduino Due has been added in 1.4.4.
Given the large number of new boards that continue to hit the market, with different pin-outs as well as boards with multiple I2C channels and based on feedback, I have decided to add the option for the user provide the communication channel also differently. This means the user will perform in the sketch the initialization of the channel (serial or I2C) and provide that to the SP30 library. Example12 (for serial communication) and Example13 (for I2C communication) has been added to demonstrate the usage.
The embedded approach as described below, and thus backward compatibility, continues to be available.
## 2.1 UART communication
### 2.1.1 ESP32
Using serial port1, setting the RX-pin(25) and TX-pin(26)
Different setup can be configured in the sketch.

SPS30 pin ESP32
1 VCC -------- VUSB
2 RX --------- TX pin 26
3 TX --------- RX pin 25
4 Select (NOT CONNECTED)
5 GND -------- GND

Also successfully tested on Serial2 (default pins TX:17, RX: 16)
NO level shifter is needed as the SPS30 is TTL 5V and LVTTL 3.3V compatible

### 2.1.2 ATMEGA2560 / Arduino Due
Used SerialPort2. No need to set/change RX or TX pin
SPS30 pin ATMEGA DUE
1 VCC -------- 5V
2 RX --------- TX2 pin 16 (DUE 17)
3 TX --------- RX2 pin 17 (DUE 18)
4 Select (NOT CONNECTED)
5 GND -------- GND

Also successfully tested on SerialPort1 and Serialport3

### 2.1.3 Failed UNO R3
Failed as I had to use softserial as there is not a separate serial port. But as the SPS30
is only working on 115K the connection failed all the time with CRC errors. (see 2.2.3)

### 2.1.4 Failed ESP8266
As the power is only 3V3 (the SPS30 needs 5V) and one has to use softserial which failed with 115K speed

## 2.2 I2C communication

As documented in the datasheet, make sure to use external 10K pull-up resistor on both the SDA and SCL lines. Otherwise the communication with the sensor will fail random. Also read chapter 4 about the expected valid values in I2C.

Although the SPS30 (according to the datasheet) can handle 100K/s, based on feedback from a user he could only the SPS30 to work for longer time stable at 50K. Example13 and example16 allow you to set the speed to 50K. (added version 1.4.10)

### 2.2.1 ESP32

SPS30 pin ESP32
1 VCC -------- VUSB
2 SDA -------- SDA (pin 21)
3 SCL -------- SCL (pin 22)
4 Select ------ GND (select I2c)
5 GND ------- GND

The pull-up resistors should be to 3V3.

Just be aware the ESP32 also has WIRE1, without default pins. You can set pins however e.g.
Wire1.begin(23,18,100000); // SDA pin 23, SCL pin 18, 100kHz frequency // Wire1 on ESP32
You still need pull-up resistors to 3V3
### 2.2.2 ATMEGA2560 / Arduino Due
SPS30 pin ATMEGA
1 VCC -------- 5V
2 SDA -------- SDA
3 SCL --------- SCL
4 Select ------- GND (select I2c)
5 GND -------- GND

On a DUE Wire already has pull-up resistors. You can also use Wire1 on a DUE. Then connect to SDA1 and SCL1 next to AREF. Here you need to add external pull-up resistors as you have to do on an MEGA2560.
### 2.2.3 UNO R3
SPS30 pin UNO
1 VCC ------- 5V
2 SDA ------- A4
3 SCL -------- A5
4 Select ------ GND (select I2c)
5 GND ------- GND

When UNO-board is detected the UART code is excluded as that does not work on UNO and will save memory. Also some buffers reduced and the call to GetErrDescription() is removed to allow enough memory. Apart from that ALL examples work with UNO as well :-)

### 2.2.4 ESP8266
SPS30 pin External ESP8266
1 VCC -------- 5V
2 SDA ----------------------- SDA
3 SCL ------------------------ SCL
4 Select ---------------------- GND (select I2c)
5 GND ------- GND -------- GND

The pull-up resistors should be to 3V3 from the ESP8266.

### 2.2.5 SODAQ AFF / SARA NB-IOT
 SPS30 pin SODAQ
 1 VCC -------- 5V
 2 SDA -------- SDA (next to ARF NOT A4 / SDA1)
 3 SCL -------- SCL (next to SCL/ARF NOT A5 / SCL1)
 4 Select ------- GND (select I2c)
 5 GND -------- GND

When connecting as indicated above to SDA and SCL those pull-up resistors are already
on the SODAQ board. SDA1 and SCL1 do NOT have pull-up and require driver change.
# 3. Software
This software has been written from scratch. The driver has no dependencies on other libraries, some examples have, due to extra sensors / hardware. It has a number of user level calls to set and get data from the SPS30 and can be used independent of the communication channel setting. The I2C communication however could have some restrictions. See chapter 4.
## 3.1 Examples overview
At the top of each sketch there is detailed information about the hardware connections and dependencies on external libraries. For a SODAQ board, the standard console output Serial (e.g. Serial.print()), needs to be changed in each example to SerialUSB.
### 3.1.1 Basic reading
This sketch will detect an SPS30 and read and display all the data every 3 seconds
It will allow to configure for communication channel, serial-RX and TX pins as well as debug level.

### 3.1.2 Basic reading with fan-clean
This sketch will detect an SPS30, read the current auto clean interval, optionally set new auto clean interval or force a manual fan-clean. After that it wile read and display all the data every 3 seconds.
It will allow to configure for communication channel, serial-RX and TX pins, debug level and options to set the auto clean interval or force a manual fan-clean.

### 3.1.3 Basic reading specific data & order
This sketch will detect an SPS30 and allows to set and display specific data in a specific order every 3 seconds. It will allow to configure for communication channel, serial-RX and TX pins, debug level and define the wanted data in a wanted order.

### 3.1.4 Basic reading with DS18X20
This sketch will detect an SPS30 and DS18X20 temperature sensor and read and display all the data every 3 seconds. It will allow to configure for the SPS30 communication channel, serial-RX and TX pins as well as debug level. For the DS18x20 it will allow to configure the pin and whether to display in Celsius or Fahrenheit. This sketch the download library: https://github.com/PaulStoffregen/OneWire. Code for the DS18x20 is based on the DS18x20 example in the onewire library
### 3.1.5 Basic reading with BME280
This sketch will detect an SPS30 and DS18X20 temperature sensor and read and display all the data every 3 seconds. It will allow to configure for the SPS30 communication channel, serial-RX and TX pins as well as debug level. For the BME280 it will allow to configure the I2C address, whether to display in Celsius or Fahrenheit and whether to display altitude in meter or foot. This sketch the download library: https://github.com/sparkfun/SparkFun_BME280_Arduino_Library

### 3.1.6 SPS30 plotter
This sketch will detect an SPS30, reads and displays the data every seconds The output format is aimed for the plotter (Cntrl+shift+L). It will allow to configure for the SPS30 communication channel, serial-RX and TX pins as well as debug level. You can configure the data to plot (MassPMx, NumPMx or average size) as well as how of the first readings to skip to remove invalid start-up info.
### 3.1.7 SPS30 and SCD30
This sketch will detect an SPS30 and SCD30, reads and display data every 3 seconds. It will allow to configure for the SPS30 communication channel, serial-RX and TX pins as well as debug level. For the SCD30 it will allow to configure to display in Celsius or Fahrenheit. This sketch the depending on the SCD30 library: https://github.com/paulvha/scd30
NOTE: on an ESP32 together with and SCD30 the standard I2C/Wire library can not be used as that does not support clock stretching. The SoftWire library, part of SCD30 has been included to support that. See 3.3.5. Apart from that the I2C on an ESP32 initializes in a particular way and as part of that it will fail often if a device is constant sending data. Read the top of the sketch for more details.
### 3.1.8 SPS30 and SCD30 and BME280
This sketch is a combination of example 5 and 7 in one sketch.
### 3.1.9 SPS30 higher precision + typical value compare
Read the SPS30 information and display with high precision. Calculated average size and compare that against the reported typical value. Provide an actual and longer running error margin.
### 3.1.10 SPS30 and AirQualityindex (AQI)
This sketch will connect to an SPS30 for getting data, store data and display the available data. The sketch has been developed and tested only for an ESP32. The data during-hour and during-day is stored RAM and at end-of-day it will be combined with data from previous days that is stored in NVRAM. Based on the captured data and the region setting, it can calculate and display an air quality index that is applicable for that region. Requires : https://github.com/paulvha/airquality
### 3.1.11 basic reading + status register
In March 2020 an updated datasheet was made available in which new SPS30 functions were indicated. These functions depend mostly on newer level firmware of the SPS30. For the status register firmware level 2.2 in the SPS30 is necessary. This example was added to show how to use the new function GetStatusReg(). The fan needs to be turning couple seconds before detecting an issue.

### 3.1.12 basic reading + any Serial
Added the option for the user to select any serial communication channel. It must be initialized in the sketch ( e.g. Serial2.begin() and then passed on the library ( sps30.begin(&Serial2)). Example12 has been added to demonstrate the usage. This “any serial channel” approach can be applied to any of the other examples.
### 3.1.13 basic reading + any I2C
Added the option for the user to select any I2C communication channel. Many of the new boards have multiple I2C channels. It must be initialized in the sketch ( e.g. Wire1.begin() and then passed on the library ( sps30.begin(&Wire1)). Example13 has been added to demonstrate the usage. This “any I2C channel” approach can be applied to any of the other examples.
### 3.1.14 basic reading + sleep and wakeup
In March 2020 an updated datasheet was made available in which new SPS30 functions were indicated. These functions depend mostly on newer level firmware of the SPS30. For the sleep and wakeup firmware level 2.0 in the SPS30 is necessary. This example was added to show how to use the new functions : sleep(), wakeup().
### 3.1.15 basic reading + any Serial + LCD
### 3.1.16 basic reading + any I2C + LCD
As Example 12 and Example 13 but added an LCD for display to display PM1, PM2.5 and PM10.
The normal background color can be set to White, Green, Red, blue or OFF.
A trigger level for PM2 can be set :
1. SET the LCD background RED if above and return to normal when below.
2. OR only turn the LCD on RED if above the limit.
A button can be defined to turn the display on for 10 seconds. This button can be combined with with option A above.
The LCD used is from Sparkfun https://www.sparkfun.com/products/16397
Needs the driver to be installed as well, see sketch header for more information.
### 3.1.20 Multiple SPS30 on the same board
This has been added in August 2020 and for more information see chapter 9
## 3.2 Overview of user level calls
### 3.2.1 EnableDebugging(act, SelectDebugSerial)
Will enable or disabled driver debug messages. Valid options for ‘act’ as level of debug
 0 : no debug message (default)
1 : sending and receiving data
2 : 1 + protocol messages

SelectDebugSerial is optional, by default it will use ‘Serial’ ( STANDARD). Starting revision 1.3.10 one could also select ‘SODAQ”, which will use ‘SerialUSB’ for debug output. See chapter 3.3.7.

### 3.2.2 SetSerialPin( rx, tx)
Sets the RX and TX pins to use for softserial or serialport1 on the ESP32. If these communication channels are used the pins must be set before calling begin. See NOTE 2 with chapter 3.2.3

### 3.2.3 begin(Comms_port)
This will initialize embedded the communication channel. Unlike in other sketches, you should NOT first perform a begin ( Serial1.begin()) in the sketch. That is all handled by the driver. The following options are supported :

I2C_COMMS	I2C communication (Wire)
SOFTWARE_SERIAL	Arduino variants (see NOTE)
SERIALPORT	ONLY IF there is NO monitor attached
SERIALPORT1	Arduino MEGA2560 and Arduino Due
ESP32 Thing : MUST define pins as defaults are for flash memory
SERIALPORT2	Arduino MEGA2560, Due and ESP32	(default is omitted)
SERIALPORT3	Arduino MEGA2560 and Arduino Due only for now

NOTE 1 : Softserial has been left in as an option, but as the SPS30 is only working on 115K the connection will probably NOT work on any device. See chapter 3.3.4 to excluding Software Serial code. On a SAM21D and Due SoftSerial is disabled.

NOTE 2: SerialPort1 is only enabled in case of MEGA2560, Arduino Due or ESP32 (with defined RX/TX before)
On some boards/IDE (like Particle Photon) Serial1 is available as well. To force-enable from the user sketch : set TX and RX pins BOTH to 8, before calling begin().

Will return ‘true’ if successful, else ‘false’.

### 3.2.4 begin (&Serial)
The library will take any initialized serial communication channel. You MUST first perform a begin ( e.g. Serial1.begin()) in the sketch, before passing on.

### 3.2.5 begin (&Wire)
The library will take any initialized Wire communication channel. (e.g. Wire1) You MUST first perform a begin (e.g. Wire1.begin()) in the sketch, before passing on.

### 3.2.6 probe()
This will try to connect the SPS30 over the requested communication channel.
Will return ‘true’ if successful, else ‘false’.

### 3.2.7 reset()
This will sent a soft-reset to the SPS30.
Will return ‘true’ if successful, else ‘false’.

### 3.2.8 start()
This will sent a start-measurement command to the SPS30.
Will return ‘true’ if successful, else ‘false’.

### 3.2.9 stop()
This will sent a stop-measurement command to the SPS30.
Will return ‘true’ if successful, else ‘false’.

### 3.2.10 clean()
This will sent a manual clean fan command to the SPS30. The fan will spin faster for 10 seconds.
The SPS30 must be in measurement mode. Will return ‘true’ if successful, else ‘false’.

### 3.2.11 GetAutoCleanInt(uint32_t *val)
This will read the current Auto Fan clean interval from the SPS30 and return that in pointer to ‘val’.
Will return ‘ERR_OK’ if successful read, else on the error messages documentation in Appendix A.

### 3.2.12 SetAutoCleanInt(uint32_t val)
This will SET the Auto Fan cleaning interval on the SPS30 to the value defined in ‘val’. Will return ‘ERR_OK’ if successful read, else on the error messages documentation in Appendix A.
### 3.2.13 GetErrDescription(code, *buf, len)
This will provide a description of an error code as documented in Appendix A. ‘Code’ the error code, ‘buf’ is a buffer to store the message and ‘len’ is the length of the buffer. (not supported on Uno : see chapter 3.2)
### 3.2.14 GetSerialNumber(*buf, len)
This will try to read the serial number from the device and store that in buffer ‘buf’ with a maximum number of bytes ‘len’. The maximum can be 32 bytes. Will return ‘true’ if successful, else ‘false’.
### 3.2.15 GetArticleCode(*buf, len)
This will try to read the Articlecode from the device and store that in buffer ‘buf’ with a maximum number of bytes ‘len’. The maximum can be 32 bytes. Will return ‘true’ if successful, else ‘false’.
!! GetArticleCode has been depreciated and for backward compatibility it will also be empty!!!
### 3.2.16 GetProductName(*buf, len)
Tries to read the product name from the device and store that in buffer ‘buf’ with a maximum number of bytes ‘len’. The maximum can be 32 bytes. Will return ‘true’ if successful, else ‘false’.
The name will always be always “00080000” , recommended to use as product identifier.
### 3.2.17 I2C_expect()
Will return the number of valid float values to expect from reading in I2C. See chapter 4.

### 3.2.18 GetValues(struct sps_values *v)
This will try to read all the values from the SPS30 and store those in structure pointer v. Will return ‘ERR_OK’ if successful read else on the error messages documentation in Appendix A.
The structure is defined in sps30.h, for details read the datasheet
typedef struct sps_values
{
    float MassPM1; // Mass Concentration PM1.0 [μg/m3]
    float MassPM2; // Mass Concentration PM2.5 [μg/m3]
    float MassPM4; // Mass Concentration PM4.0 [μg/m3]
    float MassPM10; // Mass Concentration PM10 [μg/m3]
    float NumPM0; // Number Concentration PM0.5 [#/cm3]
    float NumPM1; // Number Concentration PM1.0 [#/cm3]
    float NumPM2; // Number Concentration PM2.5 [#/cm3]
    float NumPM4; / / Number Concentration PM4.0 [#/cm3]
    float NumPM10; // Number Concentration PM4.0 [#/cm3]
    float PartSize; // Typical Particle Size [μm]
};

NOTE 1: As mentioned in chapter 4, when the communication channel is I2C, depending on the receive buffer size in wire.h, only the MassPMx may be available. The others will be zero.
NOTE 2: Although the SPS30 can provide the different values as unsigned 16 bits, this library is only reading and providing the float values. Why settle for less ?
### 3.2.19 GetSpecific value()
This will try to return a specific value as a float. In case of error a ‘-1’ value will be returned.
The following calls are valid :
GetMassPM1()
GetMassPM2()
GetMassPM4()
GetMassPM10()
GetNumPM0()
GetNumPM1()
GetNumPM2()
GetNumPM4()
GetNumPM10()
GetPartSize()
The driver will perform a GetValues() and report the specific requested value. If another value is requested this will be provided from cache. As such the different values are a snap-shot taken at the same time. If a value is requested that has been reported already all the values will be refreshed.
### 3.2.20 sleep()
This functions requires minimum firmware level 2.0. It will set the SPS30 to sleep, using even less power < 50uA according to datasheet. Measurements done indicate that the current is around 30uA
### 3.2.21 wakeup()
This functions requires minimum firmware level 2.0. It will wakeup the SPS30 after it was set to sleep. If the SPS30 was in measurement mode before turning to sleep, wakeup will restart measurement mode, else it will wakeup in idle mode. In Idle mode the SPS30 is using 320uA, in measurement mode around 52mA. a
### 3.2.22 GetVersion()
This function will read the firmware and hardware levels from the SPS30 and provide the library version level.
struct SPS30_version {
    uint8_t major;	SPS30 Firmware levels
    uint8_t minor;
    uint8_t HW_version; SPS30 hardware level (zero on I2C)
    uint8_t SHDLC_major; SHDLC protocol (zero on I2C)
    uint8_t SHDLC_minor;
    uint8_t DRV_major;	The library version level
    uint8_t DRV_minor;
};
### 3.2.23 GetStatusReg()
This functions requires minimum firmware level 2.2. This function will read the device status register. The status return is an “orred” value as multiple errors might happen at the same time.
    STATUS_OK = 0
    STATUS_SPEED_ERROR = 1 (0b00000001)
    STATUS_LASER_ERROR = 2 (0b00000010)
    STATUS_FAN_ERROR = 4 (0b00000100)
The call will return ERR_OK if NO status issues are found or ERR_OUTOFRANGE in case status issues were detected. The status-issues are cleared after reading. For fan issues to be detected the fan must be up and running for a couple of seconds.

## 3.3 Driver configuration
In the SPS30.h configuration settings are available.
### 3.3.1 Automatic configuration
In case of an UNO or another board with small memory some changes are applied automatically.
The buffers are reduced, the error message lookup is disabled and the UART communication code is excluded.
### 3.3.2 Exclude I2C communication code
In case you do not plan to use the I2C code and want to remove that in order to reduce the memory size of the driver in sps30.h comment out the line : #define INCLUDE_I2C 1.
Compile will fail if UART communication is excluded at the same time.
### 3.3.3 Exclude UART communication code
In case you do not plan to use the I2C code and want to remove that in order to reduce the memory size of the driver in sps30.h comment out the line : #define INCLUDE_UART 1
Compile will fail if I2C communication is excluded at the same time.
### 3.3.4 Exclude Software Serial communication code
In case software serial is not available for your board / IDE (causing compile errors) OR you want to save memory in sps30.h comment out the line: #define INCLUDE_SOFTWARE_SERIAL 1.
Software Serial has been in as an option, but given the 115K speed, it is expected to fail.
As SAMD device has 6 SERCOM or for the Arduino Due and Apollo3, there is no SoftSerial available. This option is to exclude.
### 3.3.5 ESP32 / SCD30 / clock stretching library
In case a sketch is to interact with both SPS30 and SCD30 on a ESP32 platform running BOTH over I2C, you MUST use the SoftWire I2C that is part of SCD30. For that you need to un-comment in sps30.h the line: // #define SOFTI2C_ESP32 1
If you platform is ESP32 and you do NOT plan to run the SPS30 over I2C, you can also apply option 3.3.2.
### 3.3.6 RX_DELAY_MS
Due to some users reporting timing issues during reading with newer SPS30 a delay has been added between written the register and reading the content. This delay is defined in SPS30.h as 100mS.
### 3.3.7 SelectDebugSerial
Starting version 1.3.10 a different serial port can be selected for debug output. For backward compatibility ‘Serial’ is choosing as default (STANDARD) with the call EnableDebugging(). An option that can be chosen with SODAQ, which will use ‘SerialUSB’ for debug output. More can be defined in the top of SPS30.h
### 3.3.8 INCLUDE_FWCHECK
New firmware levels have been slipped streamed into the SPS30 The datasheet from March 2020 shows added / updated functions on new firmware level. E.g. sleep(), wakeup(), status register are new.
On serial connection the new functions are accepted and positive acknowledged on lower level firmware, but execution does not seem to happen or should be expected.
On I2C reading Status register gives an error on lower level firmware. Sleep and wakeup are accepted and positive acknowledged on lower level firmware, but execution does not seem to happen or should be expected.
Starting version 1.4 of this library a firmware level check has been implemented and in case a function is called that requires a higher level than on the current SPS30, it will return an error.
By setting INCLUDE_FWCHECK to 0, this check can be disabled
# 4. I2C communication

In datasheet (page 16) the disclaimer is made:

Some considerations should be made about the use of the I 2 C interface. I 2 C was originally designed to connect two chips on a PCB. When the sensor is connected to the main PCB via a cable, particular attention must be paid to electromagnetic interference and crosstalk. Use as short as possible (< 10 cm) and/or well shielded connection cables.
We recommend using the UART interface instead, whenever possible: it is more robust against electromagnetic interference, especially with long connection cables

While the above is true, the I2C implementation, which works good as long as you apply the 2 x 10K pull-up resistors as indicated in the datasheet (March 2020) on pag 16.

## 4.1 Number of valid values
Opposite to an earlier assumption that the SPS30 is only sending 32 bytes, it turned out that the I2C / Wire receive buffer on many boards is set a maximum value of 32. After that the boards will stop with receiving.

A size needed for each float value is 6 bytes (LSB + MSB + CRC ) ++ (LSB + MSB + CRC) To read all values an I2C receive buffer of at least 6 x 10 = 60 bytes is needed. As on many boards the default buffer size is set to 32 in Wire.h, it will only be providing 5 valid float values.

To cope with this we determine the I2C buffer-size in the sps30.h and calculate in the constructor for sps30 the expected valid bytes. If the I2C buffer size is less than 64 only the MASS values are provided.

One can check the impact by calling I2C_expect(), which will return the number of valid float values to expect from reading SPS30.

Of course you can increase the buffer (if memory size allows) for your board in Wire.h. This is NOT part of this library, but part of your IDE.

ATMEGA (UNO or 2560 or Due .. maybe others)
I had to increase BOTH otherwise the connection would fail:
Wire/src/wire.h	#define BUFFER_LENGTH 64 // was 32
Wire/src/utility/twi.h	#define TWI_BUFFER_LENGTH 64 // was 32

Note :
Got input from Bert Heusinkveld that on a windows system for PRO-Mini (UNO variant) he had to change these values on:
-c:\program files (x86)\arduino\hardware\avr\libraries\wire…
-c:\users\...\AppData\local\arduino15\packages\arduino\hardware\avr\1.6.23\libraries\wire....

esp8266
Wire/src/wire.h	#define BUFFER_LENGTH 64 // was 32

esp32
The buffer is already 128
Wire/src/wire.h	#define I2C_BUFFER_LENGTH 128

MKRZERO / SAMD / Apollo3
The SAMD library has a different definition as it has a ‘sercom’ setup for all communication.
In wire.h for the a SAMD the RX buffer is defined as : RingBufferN<256> rxBuffer;
The library will handle that and set to 256, BUT given the 6 communication options, the SoftSerial is NOT available and thus disabled.

# 5. Relationship PM10 and PM2.5

https://www.arb.ca.gov/research/aaqs/pm/pm.htm
An PM10 surface is 16 x PM2.5, while the PM10 volume is 64 x PM2.5. This means that the <PM 2.5 as much smaller, much lighter than the PM25 <> PM10 particles. And while the PM2.5 <> PM10 numbers are much less they have a size-able impact on the numbers.
The PM10 number INCLUDE the PM2.5 numbers. Different investigations have provided information around the relationship between PM2.5 and PM10 ug/m3.
The university of Montana 70% of the weight of PM10 constitutes of PM2.5.
HongKong:
Annual(μg/m3)	Daily(μg/m3)
PM2.5 = 0.71 x PM10	PM2.5 = 0.75 x PM10

(https://www.epd.gov.hk/epd/english/environmentinhk/air/guide_ref/guide_aqa_model_g5.html)
Scotland:
The annual average ratio of daily mean PM2.5 concentration to daily mean PM10 concentration typically varied from 0.51 to 0.69 with a mean of 0.63

This was supported by the PCM modeled data, which indicated that the ratio of annual mean PM2.5 / PM10 ranged from 0.63 to 0.69

(http://www.scottishairquality.scot/assets/documents/technical%20reports/pm2.5-pm10ratio_29Mar2016-FINAL_Version_Approved.pdf)

# 6. Compare against Dylos
## 6.1 DYLOS-1700 definition
A DYLOS-1700 is providing information particle information per 0.01 cubic foot. It has 2 numbers:
A. Particle count > 0.5 μm
B. Particle count > 2.5 μm. This is a good representation of the PM10 value.(different tests)
From this you can calculate the Particle count between 0.5 μm and 2.5 μm: C = A - B
DYLOS numbers from cubic feet to cm3:	DYLOS 0.01 ft3 / 283.1685 => cm3
## 6.2 SPS30 definition
The SPS30 numbers show in cm3 the particle count SMALLER than the size. Datasheet pag 2
PMx defines particles with a size smaller than “x” micrometers (e.g., PM2.5 = particles smaller than 2.5 μm).
The SPS30 starts measurement from 0.3um
## 6.3 Comparing
Let take an example of values captured (nearly) at the same moment :
SPS30 (concentration count)
PM0.5 PM1 PM2.5 PM4 PM10
123.09985	142.61128	143.41982	143.47649	143.48987
DYLOS-1700
A = 5523 B = 42 calculated C = 5481
Compare < 2.5μm concentration
Dylos > 5481 / 283.1685 = 19.536 particles / cm3 (> 0.5)
sps30 > 143.41982 –123.09985 = 20.31 particles /cm3 (> 0.5)
Compare >2.5μm concentration
Dylos > 42 / 283.1685 = 0.148 particles /cm3
SPS30 > 143.48987 – 143.41982 = 0.07 particles /cm3
Having run the above compare 10 times, the results < 2.5μm are within 92%, where the > 2.5μm varies between 1 to 40 times difference, but on a very small amount.

Interesting / surprising to see the high number between 0.3 μm to 0.5 μm. The Dylos starts at 0.5 where the SPS40 starts at 0.3 μm.

# 7. compare against SDS-011
## 7.1 SDS-011 definition
The SDS011 using principle of laser scattering,can get the particle concentration between 0.3 to 10μm in the air. It provides the mass of 2.5 μ/m3 and 10 μ/m3 . It is not defined whether this is inclusive or exclusive of the 2.5 μm particle, but it looks they follow official definition of the < 2.5 μm particle.
## 7.2 SP30 definition
The SPS30 numbers show in cm3 the particle count SMALLER than the size. Datasheet pag 2
PMx defines particles with a size smaller than “x” micrometers (e.g., PM2.5 = particles smaller than 2.5 μm).
The SPS30 starts at 0.3 μm as well and provides mass numbers as 1 μg/m3 , 2.5 μg/m3 , 4 μg/m3 and 10 μg/m3
## 7.3 Comparing
We had the 2 sensors run side by side for ~3500 samples, every 3 seconds.

The 2.5 μ/m3 numbers are extremely close. They follow the same pattern, same reaction to change and the average over the 3500 samples is SPS30 : 4.27 μg/m3 where the SDS011 is showing : 4.45 μg/m3
Actually over a longer time it > 95% the same.

For the 10 μ/m3 the numbers are different
While they follow the same pattern and reaction, the values of SPS30 are nearly 50% of the SPS011. Over the 3500 samples on average the SPS30 is 4.35 μg/m3 where the SDS011 is 7.52 μg/m3.

We reference to chapter 5 :
Checking the relationship of SDS011 results shows over 3500 samples a factor between PM10 and PM2.5 of 0.567, often 0.6 / 0.7 drawn down especially when the μg/m3 reading is low.
Checking the relationship of SPS30 results shows over 3500 samples a factor between PM10 and PM 2.5 of 0.967, often 0.99 or 100% with lowest around 0.92.
Next to that when using example 9, compare Typical Particle Size (TPS) against average (AVG), it seems that the particles numbers above 2.5 μm to 10 μm are below of what to expect.
Using Example 9 and looking at the reported 10um mass, they are sometimes LESS than the 4um (nearly or the same as 2.5um). Also sometimes the 4 μm mass is less than 2.5 μm mass. I have 2 SPS30 sensor and BOTH have similar behavior. I have triple checked the real HEX values received from the SPS30. Below example with UART communication. :
sending: 0x7E 0x00 0x03 0x00 0xFC 0x7E
Received: 0x7E 0x00 0x03 0x00 0x28 0x41 0xB5 0x07 0x06 0x41 0xBF 0x36 0x47 0x41 0xBF 0x36 0x65 0x41 0xBF 0x36 0x47 0x43 0x18 0x06 0xDA 0x43 0x30 0xFB 0xD9 0x43 0x31 0xA0 0x35 0x43 0x31 0xA3 0x50 0x43 0x31 0xA4 0xB7 0x3F 0x09 0xD2 0x78 0xA9 0x7E length: 46
0x7E	header
0x00	address
0x03	command
0x00	state
0x28	length
0x41 0xB5 0x07 0x06	1 ug/m3
0x41 0xBF 0x36 0x47	2.5 ug/m3
0x41 0xBF 0x36 0x65	4 ug/m3
0x41 0xBF 0x36 0x47	10 ug/m3 (PM10 mass values is the same as PM.25 and lower than PM4 ?)
0x43 0x18 0x06 0xDA	0.5 #/m3
0x43 0x30 0xFB 0xD9	1 #/m3
0x43 0x31 0xA0 0x35	2.5 #/m3
0x43 0x31 0xA3 0x50	4 #/m3
0x43 0x31 0xA4 0xB7	10 #/m3 (but at the same time the PM10 particle count is higher ???)
0x3F 0x09 0xD2 0x78	Typical particle
0xA9	CRC
0x7E	trailer
# 8 Conclusions
## 8.1 Personal conclusion (feb 2019)
I believe the mass and volume numbers up to 2.5um are well representing the environment. But I wonder whether the sensor is able to detect / report above 2,5 μm particle size the average weight and/or the number of particles in the right way.
## 8.2 Additional / external PMx remarks
Since the release of the SPS30 a number of people and instances have tested this device and the remarks about PM10 measurement have been the same : is this correct ?

Many of the people who have reached out to me came to similar conclusions and wondered.
A research paper concluded that the Sensirion Sps30 is not able to detect particles above ~2 um.
https://www.researchgate.net/publication/337852062_Laboratory_evaluation_of_particle_size-selectivity_of_optical_low-cost_particulate_matter_sensors
Document available in extra folder : Laboratory_evaluation_of_particle_size-selectivity
Sensirion has been providing additional information. A specialist article has been created (copied in the extras directory).
“The quality of the mass estimation will thus vary highly depending on the manufacturer algorithm used to convert the measured optical signal into PM mass concentration…….
Sensirion’s proprietary algorithms use an advanced approach that allows a proper estimation of the mass concentration, regardless of the particle type measured. In addition, such an approach enables a correct estimation of the size bins. Also, an additional bin output is provided in contrast to most state-of-the-art consumer PM sensors on the market – PM4.0.”
The datasheet March 2020, footnote 4 on page 2:
“PM4 and PM10 output values are calculated based on distribution profile of all measured particles”

All in all it supports the earlier personal conclusions, but this additional information provide more background information.

# 9. Multiple SPS30 at the same time

Let’s say you want to measure the air quality on different places at the same time with different SPS30 :

- Inside the building different rooms
- Inside the building different rooms and outside the building
- Inside the building different floors

BUT you do not want to have multiple CPU boards !!

Because you may want to combine or compare the data on a single place or act on the measurements from a single place with an action or alarm condition….

GOOD NEWS… YOU CAN connect up to at least 5 (FIVE !) SPS30 to a single board.

Starting with example20 you can connect up to 5 (FIVE !) different SPS30 to a single board. They can be connected to either Serial1, Serial2, Serial3 or also at the same time to Wire, Wire1 or any WireX you might have. Example20 allows up to 3 Serial and up to 2 Wire connections at the same time !! With a little change, you can add Serial connection or add Wire connections (if your board support that) .

The example20 has been tested on a Arduino MEGA2560, Arduino DUE and ESP32. The MEGA2560 and ESP32 can handle 4 while Arduino DUE can have 5 running at the same time.

As I ‘only’ have 3 SPS30 I was able to test with 3 in parallel. In this chapter I will share the setup and the experience I had.

## 9.1 Hardware device setup

This is already described in chapter 2 with BOTH Serial and I2C on the different boards.
Just be reminded that for each Wire-channel you have to use different pull-up resistors !!

### 9.1.1 I2C cable length
The maximum bus length of an I2C link is about 1 meter at 100 Kbaud, or 10 meters at 10 Kbaud. All driven by the quality of the cable and it’s capacitive value. There are external I2C cable extenders with can extend to 300 meters. That said, I have not tested the library using an extender and that might need to include some delays. The default speed for I2C channel is 100Khz, BUT you could lower that by added a call SPS_COMMSx.setClock(clockFrequency), e.g. 35000 as 35Khz, for SPS30_COMMS4 or SPS30_COMMS5.

### 9.1.2 Serial / RS232 cable length

The RS-232 standard specifies this as 15 meters. On lower baud rates the lines can be longer (e.g. 100 meter at 9600 baud), but the SPS30 only works on 115K. There are external RS232 extenders which can extend to as far as 1200 meters on 115K. That said, I have not tested the library using an extender and that might need to include some delays. In the file sps30.h from the library a variable RX_DELAY_MS is defined as 100mS. This delay between requesting data and reading data could be extended if needed. The received data is normally buffered before reading by the application.

## 9.2 Software
### 9.2.1 General settings
The software is modular and has a couple of settings.
Chapter 3.3 defines the settings in sps30.h
Chapter 4 is especially important for the number of values in case of I2C

### 9.2.2 Sketch
The Example20 sketch has a number of settings

Communication channel

Around line 116 you define the number of SPS30 as well as the communication channels
Comment out ( like SPS_COMMS3) the channels that you do NOT use.

- Serial channel for SPS30_COMMS1 – SPS30_COMMS3 and / or
- I2C channel for SPS30_COMMS4 – SPS30_COMMS5

#define SPS30_COMMS1 Serial1 // serial only
#define SPS30_COMMS2 Serial2 // serial only
//#define SPS30_COMMS3 Serial3 // serial only
#define SPS30_COMMS4 Wire // I2C only
#define SPS30_COMMS5 Wire1 // I2C only

Debug

Debug messages can be set for each communication channel individually, DEBUG1 – DEBUG5.
 0 : no DEBUG messages
 1 : request sending and receiving
 2 : request sending and receiving + show protocol errors

#define DEBUG1 0
#define DEBUG2 0
#define DEBUG3 0
#define DEBUG4 0
#define DEBUG5 0

ESP32 Serial1

The ESP32 is flexible is assigning the serial communication to different pins. The default Serial1 pins are used by memory, but you can still use the Serial1 by assigning different pins. See line 147 in the sketch for Serial1 instructions : SPS30_COMMS1.begin(115200, SERIAL_8N1, 25, 26); // ESP32 in case of Serial1

ESP32 Wire1

The ESP32 has 2 Wire channels. The Wire1 does not have default pins and these need to be defined as you setup the connection with: SPS30_COMMS5.begin(23,18,100000); // SDA pin 23, SCL pin 18, 100kHz frequency // Wire1 on ESP32
## 9.3 Multiple connection test

As stated before I have 3 SPS30, but WHY ?

The first one I bought had no connector included so I could not use it. What a deception ! The second I bought from Sparkfun, these guys get it… you want to test and thus a connector is included. Both running Firmware level 1.0 The third one I got from Sensirion, as an appreciation from the work done on the library (Thank You) BUT no connector (and not expected)

At later stage I have ordered the right connectors from AliExpress and can now connect all 3 devices.

### 9.3.1 Arduino DUE
The DUE is able to support up to 5 devices: 3 x serial + 2 x Wire. In my test I connected as follows:

SPS30-1 : connection Serial1

SPS30 pin DUE
1 VCC -------- 5V
2 RX --------- TX2 pin 18
3 TX --------- RX2 pin 19
4 Select (NOT CONNECTED)
5 GND -------- GND

#define SPS30_COMMS1 Serial1 // serial only

SPS30-2 : connection Serial2

SPS30 pin DUE
1 VCC -------- 5V
2 RX --------- TX2 pin 16
3 TX --------- RX2 pin 17
4 Select (NOT CONNECTED)
5 GND -------- GND

#define SPS30_COMMS2 Serial2 // serial only

SPS30-5: connection Wire1 (with external pull-up)

SPS30 pin DUE
1 VCC -------- 5V
2 SDA -------- SDA1	NEXT TO AREF
3 SCL --------- SCL1	NEXT TO SDA1
4 Select ------- GND (select I2c)
5 GND -------- GND

#define SPS30_COMMS5 Wire1 // I2C only

Last by not least, comment out the lines:

//#define SPS30_COMMS3 Serial3 // serial only
//#define SPS30_COMMS4 Wire // I2C only

The Arduino DUE standard Wire (NOT Wire1) keeps confusing me. Sometimes the standard Wire (SDA0 / SCL0) works and sometimes it does not. This has to do with the on-board resistors, which should NOT be there – only the client / slave devices. In the tests I have done it did work, but with other slave devices with pull-up resistors it did not work. While NOT applying external pull-up resistors Wire worked. It worked always with Wire1, if you not forget to add external pull-up resistors of 5K6 or 10K.

### 9.3.2 Arduino Mega2560
The Mega2560 is able to support up to 4 devices: 3 x serial + 1 Wire. There is no Wire1 on the Arduino Mega2560, the SDA and SCL pins are just 2 times available. I connected as follows, just for a test:

SPS30-1 : connection Serial1

SPS30 pin ATMEGA
1 VCC -------- 5V
2 RX --------- TX2 pin 19
3 TX --------- RX2 pin 18
4 Select (NOT CONNECTED)
5 GND -------- GND

#define SPS30_COMMS1 Serial1 // serial only

SPS30-2 : connection Serial2

SPS30 pin ATMEGA
1 VCC -------- 5V
2 RX --------- TX2 pin 16
3 TX --------- RX2 pin 17
4 Select (NOT CONNECTED)
5 GND -------- GND

#define SPS30_COMMS2 Serial2 // serial only

SPS30-4: connection Wire

SPS30 pin ATMEGA
1 VCC -------- 5V
2 SDA -------- SDA
3 SCL --------- SCL
4 Select ------- GND (select I2c)
5 GND -------- GND

#define SPS30_COMMS4 Wire // I2C only
Last by not least, comment out the lines:

//#define SPS30_COMMS3 Serial3 // serial only
//#define SPS30_COMMS5 Wire1 // I2C only

### 9.3.3 ESP32
Is able to support up to 4 devices: 2 x serial + 2 x Wire. In my test I connected as follows:

SPS30-1 : connection Serial1

SPS30 pin ESP32
1 VCC -------- 5V
2 RX --------- TX2 pin 26
3 TX --------- RX2 pin 25
4 Select (NOT CONNECTED)
5 GND -------- GND

#define SPS30_COMMS1 Serial1 // serial only

I have then changed around line 330 from
SPS30_COMMS1.begin(115200);
to
SPS30_COMMS1.begin(115200, SERIAL_8N1, 25, 26); // ESP32 in case of Serial1

SPS30-4: connection Wire

SPS30 pin ESP32
1 VCC -------- 5V
2 SDA -------- SDA pin 21
3 SCL --------- SCL	pin 22
4 Select ------- GND (select I2c)
5 GND -------- GND

#define SPS30_COMMS4 Wire // I2C only
Pin 21 and 22 are the default pins for Wire, but could be changed like done for Wire.
Also added 2 pull-up resistors of 5k6

SPS30-5: connection Wire1

SPS30 pin ESP32
1 VCC -------- 5V
2 SDA -------- SDA pin 23
3 SCL --------- SCL	pin 18
4 Select ------- GND (select I2c)
5 GND -------- GND

#define SPS30_COMMS5 Wire1 // I2C only

I have then changed around line 449 from
SPS30_COMMS5.begin();
to
SPS30_COMMS5.begin(23,18,100000); // SDA pin 23, SCL pin 18, 100kHz frequency // Wire1 on ESP32

Also added 2 pull-up resistors of 5k6

In case you want to use Serial2 as well :

SPS30-2 : connection Serial2

SPS30 pin ESP32
1 VCC -------- 5V
2 RX --------- TX2 pin 17
3 TX --------- RX2 pin 16
4 Select (NOT CONNECTED)
5 GND -------- GND

#define SPS30_COMMS2 Serial2 // serial only
Pin 16 and 17 are the default pins for Serial2, but could be changed like done for Serial1

## 9.4 Sample output ESP32
SPS30-Example20: Basic reading with multiple SPS30. press <enter> to start
Trying to connect.

Detected SPS30 - 1
Serial number : 361D4FD3AA941A76 Product name : 00080000
Firmware level: 1.0 Hardware level: 6 SHDLC protocol: 2.0 Library level : 1.4
Measurement started SPS30 - 1

Detected SPS30 - 4
Serial number : 1FDFF1525C39D2B1 Product name : 00080000
Firmware level: 1.0 Hardware level: 0
Measurement started SPS30 - 4

Detected SPS30 - 5
Serial number : 46B997159016E02F Product name : 00080000
Firmware level: 2.2 Hardware level: 0
Measurement started SPS30 - 5

Hit <enter> to continue reading.

SPS30 -------------Mass ----------- ------------- Number -------------- -Average- --Serialnumber--
Dev Concentration [μg/m3] Concentration [#/cm3] [μm]
      P1.0	P2.5	P4.0	P10	P0.5	P1.0	P2.5	P4.0	P10	PartSize

1 3.11	3.61	3.92	4.03	20.60	24.48	24.96	25.01	25.02	0.62	361D4FD3AA941A76
4 5.63	10.23	13.61	14.46	27.20	39.61	44.40	45.00	45.10	0.85	1FDFF1525C39D2B1
5 6.14	6.49	6.49	6.49	41.85	48.78	48.99	49.01	49.02	0.52	46B997159016E02F

1 3.77	4.36	4.73	4.86	25.00	29.69	30.26	30.32	30.34	0.63	361D4FD3AA941A76
4 5.61	9.95	13.13	13.93	27.64	39.65	44.18	44.74	44.84	0.87	1FDFF1525C39D2B1
5 6.02	6.36	6.36	6.36	41.04	47.83	48.04	48.06	48.07	0.49	46B997159016E02F

1 4.85	5.53	5.94	6.08	32.41	38.32	38.96	39.03	39.05	0.60	361D4FD3AA941A76
4 4.94	7.66	9.60	10.08	27.11	36.10	38.89	39.23	39.29	0.74	1FDFF1525C39D2B1
5 6.04	6.39	6.39	6.39	41.17	47.99	48.19	48.21	48.23	0.53	46B997159016E02F

1 5.52	6.21	6.61	6.75	37.11	43.71	44.36	44.43	44.44	0.60	361D4FD3AA941A76
4 5.36	7.96	9.78	10.24	30.32	39.56	42.20	42.53	42.58	0.74	1FDFF1525C39D2B1
5 6.19	6.54	6.54	6.54	42.18	49.15	49.36	49.39	49.40	0.51	46B997159016E02F

1 5.86	6.57	6.97	7.12	39.45	46.42	47.09	47.16	47.17	0.62	361D4FD3AA941A76
4 6.07	8.77	10.64	11.11	34.92	45.03	47.76	48.09	48.15	0.73	1FDFF1525C39D2B1
5 6.48	6.85	6.85	6.85	44.16	51.46	51.68	51.71	51.72	0.49	46B997159016E02F

# Appendix A : Error messages
Defined in sps30.h

Hex
00 ERR_OK	“All good"
01 ERR_DATALENGTH "Wrong data length for this command (too much or little data)"
02 ERR_UNKNOWNCMD "Unknown command"
03 ERR_ACCESSRIGHT "No access right for command"
04 ERR_PARAMETER "Illegal command parameter or parameter out of allowed range"
28 ERR_OUTOFRANGE "Internal function argument out of range"
43 ERR_CMDSTATE "Command not allowed in current state"
50 ERR_TIMEOUT "No response received within timeout period"
51 ERR_PROTOCOL "Protocol error"
88 ERR_FIRMWARE “Not supported on this SPS30 firmware level”

# Appendix B : some example output

## Example 1 : Basic reading

SPS30-Example1: Basic reading. press <enter> to start
Trying to connect
Detected SPS30.
Serial number : 361D4FD3AA941A76
Product name : 00080000
Firmware level: 1.0
Library level: 1.4
Measurement started

Hit <enter> to continue reading
Hit <enter> to continue reading (Note 1)
-------------Mass ----------- ------------- Number -------------- -Average-
     Concentration [μg/m3] Concentration [#/cm3] [μm]
P1.0	P2.5	P4.0	P10	P0.5	P1.0	P2.5	P4.0	P10	PartSize

13.31	16.19	18.14	18.85	107.15	103.87	106.74	107.09	0.00	0.60
12.88	14.86	16.06	16.51	103.75	101.60	103.49	103.71	0.00	0.61
12.77	14.80	16.06	16.52	102.83	100.61	102.56	102.79	0.00	0.65
12.50	14.04	14.90	15.22	100.64	99.02	100.45	100.61	0.00	0.61
12.14	13.42	14.08	14.32	97.77	96.48	97.63	97.75	0.00	0.62
11.97	13.09	13.63	13.82	96.39	95.29	96.27	96.37	0.00	0.61
11.80	12.87	13.37	13.55	95.02	93.98	94.92	95.01	0.00	0.61
11.61	12.60	13.04	13.20	93.50	92.55	93.40	93.48	0.00	0.60
11.50	12.44	12.84	12.99	92.59	91.70	92.50	92.58	0.00	0.59
11.37	12.29	12.67	12.81	91.58	90.72	91.49	91.56	0.00	0.59
11.43	12.32	12.68	12.82	92.02	91.19	91.94	92.01	0.00	0.57
11.43	12.31	12.65	12.77	92.07	91.27	92.00	92.06	0.00	0.58
11.43	12.28	12.60	12.72	92.09	91.32	92.01	92.07	0.00	0.57
11.34	12.15	12.45	12.56	91.33	90.60	91.26	91.31	0.00	0.57
11.25	12.03	12.31	12.41	90.58	89.88	90.51	90.57	0.00	0.56
11.10	11.85	12.10	12.19	89.42	88.77	89.37	89.41	0.00	0.56
11.13	11.85	12.08	12.17	89.61	88.98	89.55	89.60	0.00	0.56
11.00	11.69	11.90	11.98	88.56	87.97	88.51	88.55	0.00	0.57
10.76	11.42	11.61	11.68	86.67	86.12	86.63	86.66	0.00	0.56
10.71	11.35	11.53	11.60	86.23	85.70	86.19	86.23	0.00	0.56
10.67	11.30	11.47	11.53	85.94	85.42	85.90	85.93	0.00	0.56
10.56	11.17	11.33	11.39	85.02	84.52	84.98	85.01	0.00	0.57
10.51	11.11	11.26	11.32	84.62	84.13	84.58	84.61	0.00	0.56
10.55	11.16	11.31	11.36	85.00	84.52	84.97	85.00	0.00	0.57
10.73	11.36	11.51	11.57	86.46	85.96	86.42	86.45	0.00	0.57
11.08	11.72	11.88	11.94	89.23	88.72	89.19	89.22	0.00	0.56
11.17	11.81	11.96	12.02	89.95	89.44	89.91	89.94	0.00	0.56
11.01	11.63	11.78	11.83	88.66	88.16	88.62	88.65	0.00	0.57
10.89	11.50	11.65	11.70	87.70	87.22	87.67	87.70	0.00	0.58

## Example 5 : Basic reading with BME280
SPS30-Example5: Basic reading + BME280. press <enter> to start
Trying to connect
Detected SPS30
Serial number : 361D4FD3AA941A76
Product name : 00080000
Firmware level: 1.0
Library level: 1.4
Detected BME280
Measurement started

Hit <enter> to continue reading (Note 1)
================================= SPS30 ================================= ========== BME280 ============
--------------------Mass ------------------- ----------------------- Number ------------------------- -Average- Pressure Humidity Altitude Temperature
        Concentration [μg/m3] Concentration [#/cm3] [μm] [hPa]	[%] Meter	[*C]
P1.0	P2.5	P4.0	P10	P0.5	P1.0	P2.5	P4.0	P10 PrtSize

15.26	15.87	15.87	15.87	122.90	122.53	122.90	122.90	0.00	0.36	967	37.7	390.6	22.98
15.69	16.33	16.33	16.33	126.39	126.01	126.38	126.39	0.00	0.42	1005	37.8	68.8	22.93
15.90	16.54	16.54	16.54	128.06	127.67	128.05	128.06	0.00	0.44	1005	37.8	68.9	22.90
16.63	17.30	17.30	17.30	133.93	133.52	133.92	133.93	0.00	0.45	1005	37.7	67.8	22.84
17.01	17.70	17.70	17.70	137.02	136.60	137.01	137.01	0.00	0.46	1005	37.6	69.1	22.80
17.20	17.89	17.89	17.89	138.51	138.08	138.50	138.50	0.00	0.46	1005	37.5	68.5	22.77
17.20	17.89	17.89	17.89	138.51	138.09	138.50	138.51	0.00	0.47	1005	37.6	69.4	22.73
17.24	17.94	17.94	17.94	138.89	138.47	138.88	138.89	0.00	0.47	1005	37.5	68.9	22.67
17.40	18.10	18.10	18.10	140.13	139.70	140.12	140.13	0.00	0.47	1005	37.5	68.4	22.63
17.42	18.12	18.12	18.12	140.30	139.88	140.30	140.30	0.00	0.48	1005	37.3	68.6	22.58

## Example 7 : Basic reading with SCD30

SPS30-Example7: Basic reading + SCD30. press <enter> to start
Trying to connect
Serial number : 361D4FD3AA941A76
Product name : 00080000
Firmware level: 1.0
Library level: 1.4
SCD30 found
Serial number: 390943
Measurement started

Hit <enter> to continue reading	(Note 1)
=================================== SPS30 ================================= ========= SCD30 ==========
---------------------Mass ----===---------- ----------------------- Number -------------------------- -Average- CO2 Humidity Temperature
     Concentration [μg/m3] Concentration [#/cm3] [μm] [ppm] [%] [*C]
P1.0	P2.5	P4.0	P10	P0.5	P1.0	P2.5	P4.0	P10 PrtSize

16.29	17.21	17.21	17.21	127.91	127.44	127.90	127.91	0.00	0.55	741	35.6	22.15
17.86	21.66	23.86	24.41	140.80	136.81	140.33	140.73	0.00	0.57	741	35.4	22.25
17.08	19.54	20.73	21.02	134.41	132.05	134.16	134.37	0.00	0.56	742	35.3	22.33
15.96	16.86	16.86	16.86	125.31	124.85	125.30	125.31	0.00	0.54	742	35.0	22.44
15.65	16.53	16.53	16.53	122.85	122.40	122.84	122.85	0.00	0.55	741	35.0	22.49
15.64	17.00	17.37	17.46	122.90	121.85	122.81	122.88	0.00	0.56	739	34.8	22.62
15.61	16.49	16.49	16.49	122.56	122.11	122.55	122.56	0.00	0.56	739	34.7	22.67
15.97	16.89	16.90	16.91	125.39	124.90	125.37	125.38	0.00	0.56	744	34.5	22.77
16.12	17.05	17.08	17.08	126.54	126.03	126.52	126.53	0.00	0.56	745	34.4	22.80

## Example 8 : Basic reading + SCD30 + BME280
Trying to connect
Detected SPS30
  Serial number : 361D4FD3AA941A76
  Product name : 00080000
  Firmware level: 1.0
  Library level: 1.4
Detected SCD30
    Serial number : 390943
Detected BME280
Measurement started

Hit <enter> to continue reading (Note 1)

============== SPS30 ============================== ========= SCD30 ====== ======== BME280 ===============
-------------Mass ----------- ------------- Number -------------- -Average- CO2 Humidity Temperature Pressure Humidity Altitude Temperature
     Concentration [μg/m3] Concentration [#/cm3] [μm] [ppm] [%] [*C] [hPa] [%] Meter [*C]
P1.0 P2.5 P4.0 P10 P0.5 P1.0 P2.5 P4.0 P10 PrtSize

3.04 3.63 3.97 4.05 23.96 23.34 23.88 23.94 0.00 0.60 804 32.7 24.89 1022 44.7 -74.1 21.29
2.91 3.62 4.05 4.15 22.98 22.22 22.89 22.97 0.00 0.59 802 32.6 24.89 1022 44.6 -74.0 21.29
2.57 3.05 3.31 3.38 20.25 19.76 20.19 20.24 0.00 0.60 800 32.6 24.90 1022 44.8 -74.0 21.30
2.35 2.82 3.09 3.16 18.54 18.05 18.48 18.53 0.00 0.60 799 32.5 24.87 1022 45.3 -74.8 21.30
2.18 2.50 2.66 2.69 17.13 16.81 17.09 17.12 0.00 0.56 796 32.6 24.89 1022 45.0 -74.6 21.30
2.15 2.57 2.81 2.87 16.95 16.52 16.90 16.95 0.00 0.57 797 32.7 24.87 1022 44.6 -74.4 21.30
2.01 2.34 2.52 2.56 15.82 15.48 15.78 15.81 0.00 0.57 798 32.7 24.90 1022 45.4 -74.1 21.30
1.95 2.32 2.53 2.58 15.39 15.01 15.35 15.39 0.00 0.60 798 32.8 24.89 1022 45.1 -74.5 21.31
1.84 2.21 2.42 2.47 14.48 14.10 14.44 14.48 0.00 0.58 800 32.7 24.87 1022 45.0 -74.2 21.30
1.76 2.07 2.24 2.28 13.90 13.58 13.86 13.89 0.00 0.58 801 32.7 24.89 1022 45.0 -75.1 21.30

## Example 9 : Basic reading + typical + average
Hit <enter> to continue reading
-----------------------Mass --------------------- -------------------------- Number --------------------------- -----Partsize -------- ----- Error Margin -----
        Concentration [μg/m3] Concentration [#/cm3] [μm] %
 PM1.0 PM2.5 PM4.0 PM10 PM0.5 PM1.0 PM2.5 PM4.0 PM10 Typical Average Actual last 10 samples

 2.44646	3.72065	4.61890	4.84432	13.60713 17.95112 19.25093 19.41132 19.43798 0.68469	0.78369	-14.46	-14.46
 2.98158	3.99828	4.66924	4.83762	17.91819 22.43591 23.43956 23.55998 23.58018 0.66387	0.70377	-6.01	-10.30
 3.25240	4.01541	4.47382	4.58886	20.40710 24.83409 25.55428 25.63722 25.65131 0.62317	0.65614	-5.29	-8.72
 3.52776	4.20484	4.58310	4.67803	22.50948 27.09331 27.71152 27.78041 27.79225 0.58633	0.63696	-8.64	-8.70
 3.68784	4.34082	4.69292	4.78127	23.66743 28.37988 28.96679 29.03113 29.04225 0.56174	0.63027	-12.20	-9.33
 3.82785	4.54891	4.94859	5.04888	24.45823 29.41227 30.06833 30.14116 30.15369 0.59167	0.63536	-7.39	-9.02
 3.90359	4.63237	5.03476	5.13575	24.95849 30.00106 30.66302 30.73637 30.74899 0.59676	0.63461	-6.34	-8.65
 3.94991	4.69951	5.11631	5.22090	25.22431 30.34434 31.02732 31.10326 31.11631 0.60294	0.63599	-5.48	-8.26
 3.94206	4.61696	4.97509	5.06495	25.35640 30.36026 30.96264 31.02818 31.03953 0.59949	0.62764	-4.70	-7.87
 3.85170	4.45136	4.75403	4.82998	24.92396 29.72656 30.25042 30.30610 30.31580 0.60135	0.62064	-3.21	-7.41
 3.79606	4.38160	4.67558	4.74935	24.57754 29.30286 29.81325 29.86734 29.87679 0.60597	0.61999	-2.31	-6.52
 3.74854	4.33460	4.63112	4.70552	24.25030 28.92784 29.44034 29.49486 29.50437 0.60642	0.62094	-2.39	-6.11
 3.73187	4.26205	4.51513	4.57863	24.27512 28.85470 29.30724 29.35406 29.36230 0.60149	0.61450	-2.16	-5.79
 3.81722	4.27159	4.46097	4.50849	25.04913 29.60613 29.97384 30.00942 30.01584 0.58703	0.60410	-2.91	-5.07
 3.86197	4.29266	4.46133	4.50366	25.41500 29.98342 30.32403 30.35597 30.36181 0.58214	0.60071	-3.19	-4.09
 3.84442	4.25367	4.40619	4.44445	25.34799 29.86744 30.18542 30.21449 30.21984 0.58537	0.59842	-2.23	-3.63
 3.93010	4.40392	4.60364	4.65376	25.77495 30.47541 30.86048 30.89796 30.90471 0.60476	0.60479	-0.01	-3.18
 3.79063	4.30330	4.53995	4.59933	24.72161 29.33589 29.76757 29.81152 29.81931 0.61342	0.61143	0.32	-1.61
 3.70268	4.20668	4.44037	4.49902	24.14001 28.65187 29.07703 29.12041 29.12809 0.61379	0.61182	0.32	-1.41

Note 1:
The PM10 showing 0 in some examples was an error that has been fixed in the code in a later version
# Appendix C: inside SPS30

Removing the top cover screw

 Carefully remove from the housing. Watch the spring at the back. Removing the 4 screws allows the top panel to be removed

Air goes in on the right hand side, through a measurement chamber and is coming out through the fan. At the bottom, the golden piece, is the led.